PE2
        U5
GND 3V3 5V
   | | |
PE0 PE1 PE2 PE3 PE4 PE5
## Jitter-free Calibration Out for HANTEK 6022BE/BL

An issue of the Hantek scope is the jitter of the calibration out signal. It is generated by toggling an output pin in an interrupt routine triggered by counter/timer 2 overflow.
Due to interrupt latency the signal jitters by some µs, this leads to a high phase jitter on high output frequencies. The EzUSB can provide the T2 overflow signal on pin PE2/T2OUT (PE2, available on JP2 of the PCB, at least 6022BE) as an alternate port function. According to TRM (rev. G, p.193) the signal is "one CLKOUT pulse".
Connect a toggle flipflop to PE2 to create the correct output frequency with 50% duty cycle.
You can use e.g. 74HC74 (dual D flipflop), 74HC161 / 74HC163 (4-bit binary counter), or 74HC4024 ( 7-bit binary counter). These are even available in maker-friendly DIL packages.
Depending on your needs, supply the chip either with 5V (pin 3 of U5 - AMS1117-3.3) or with 3.3V (pin 2 of U5), GND is U5/pin1.

The 6022BL has no JP2, you must solder directly on pin 88 (PE2/T2OUT)
## Original Software Solution
Software created calibration output for 6022BE:
#define TOGGLE_CALIBRATION_PIN() \
    do { \
        PA7 = !PA7; \
    } while ( 0 )

void timer2_isr( void ) __interrupt TF2_ISR {
    /* Toggle the probe calibration pin. */
    TOGGLE_CALIBRATION_PIN();
#ifdef LED_RED_TOGGLE
    // Avoid nasty sdcc 4.0 REGRESSION:
    // Do not use "if ( ledcounter && --ledcounter == 0 )"
    // Write separate statements!
    // Otherwise the ISR uses registers for reload -> additional push/pop ...
    // ... more cycles, fails for 100 kHz
    if ( ledcounter ) {
        --ledcounter;
        if ( ledcounter == 0 ) { // led timed out?
        ledcounter = ledinit; // reload
        LED_RED_TOGGLE();
        }
    }
#endif
    TF2 = 0;
}

SDCC builds this assembler code, that toggles almost immediately (line 1209):
      000136 1205 _timer2_isr:
      000136 C0 E0 [24] 1206 push acc
      000138 C0 D0 [24] 1207 push psw
        1208 ; scope6022.inc:78: TOGGLE_CALIBRATION_PIN();
      00013A B2 87 [12] 1209 cpl _PA7
        1210 ; scope6022.inc:85:
if ( ledcounter ) {
      00013C E5*06 [12] 1211 mov a,_ledcounter
      00013E 45*07 [12] 1212 orl a,(_ledcounter + 1)
      000140 60 19 [24] 1213 jz 00110$
…

The neccessary firmware modification sets bit 2 of PORTECFG to 1, this enables the alternate function T2OUT on PE2.
/* PE2: T2OUT for HW calibration output, other bits are port E output */
#define INIT_PORTECFG 0x04